TapeStation RNA Quantification

Written by M. Studivan

Updated: 7 May 2026 M. Studivan

- 1) Allow RNA Sample Buffer (stored at 4 ºC) to come to room temperature for 30 min.

- 2) Thaw RNA Ladder (stored at -20 ºC) and RNA samples on ice.

- 3) Turn on the TapeStation and launch the Controller software.

- 4) Open a new RNA ScreenTape (or use one with lanes remaining) and flick it gently to remove any air bubbles. ScreenTapes are stored at 4 ºC.

- 5) Insert the ScreenTape into the ScreenTape nest on the instrument, with the barcode facing back on the right-hand side. The software will recognize the tape and the number of available (light gray) lanes vs used lanes (dark gray).

- 6) Select the required number of sample positions in the software and add sample IDs. The software will display the required number of consumables (tips, tube strips, tapes).

- 7) Vortex and spin down reagents.

- 8) Optional: Prepare ladder: add 5 µL RNA Sample Buffer and 1 µL RNA Ladder at position A1 in a tube strip.

- 9) Add 5 µL of buffer to the remaining tubes for each sample, then add 1 µL of RNA sample to the respective tube.

- 10) Cap the tube strip(s), then vortex on the IKA MS3 vortexer at 2000 rpm for 1 min (just load and press start). Quick spin to bring everything to the bottom.

- 11) Denature samples and ladder: heat tube strips at 72 ºC for 3 min, cool them down for 3 min, followed by a quick spin. This step is done in the thermocycler under the TapeStation RNA program.

- 12) Load tube strips into the TapeStation, with the ladder at position A1 (if used).

- 13) Remove the strip caps carefully and check that the liquid is still at the bottom.

- 14) Select Start.

- 15) The Analysis software will open and save a project file automatically when complete.

- 16) Empty the tip waste bucket and remove the tape. If the tape still has remaining lanes, put it back in 4 ºC with the reagents for up to 2 weeks. Return all reagents to proper storage temperatures.

TapeStation Consumable Prep:

- TapeStation tips (insert new rack when empty)

- TapeStation strip tubes and caps (enough for samples +1 ladder)